Calibration and Measuring Equipment Procedure

| Document No. | QMS-PRO-006 | Standard | ISO 9001:2015 Cl. 7.1.5 |
| --- | --- | --- | --- |
| Revision | 2.0 | Issue Date | [Date] |
| Review Frequency | Annual | Owner | QA Inspector / Quality Manager |

# 1. Purpose

To ensure that measuring and test equipment (M&TE) used to determine product conformance is fit for purpose, traceable to recognised measurement standards, and used and maintained in a way that gives confidence in measurement results.

# 2. Scope

Applies to all M&TE used at [Organisation Name] for conformance decisions, regardless of ownership ([Organisation Name]-owned, hired, employee-provided, customer-provided). Includes:

- Dimensional measurement: verniers, micrometers, calipers, tape measures used for acceptance, height gauges.

- Pressure equipment: gauges used for testing.

- Torque equipment: torque wrenches, torque indicators.

- NDT equipment: thickness gauges, hardness testers, magnetic flux indicators.

- Weighing equipment used for acceptance.

- Temperature equipment used for process control affecting product conformance.

Equipment used only for indication (not acceptance decisions) is identified separately and not subject to this procedure's calibration requirements.

# 3. M&TE Register (QMS-REG-CAL-001)

All M&TE in scope is recorded in the Calibration Register with: unique item number, description, make/model, serial number, location, owner, calibration interval, last calibration date, calibration due date, status.

# 4. Calibration Intervals

- Default —  12 months for most instruments.

- Shorter intervals —  Apply where the instrument is used in harsh conditions, used for high-risk acceptance, or has demonstrated drift in past calibrations.

- Longer intervals —  Up to 24 months only where multiple calibrations have shown stable performance and supported by documented analysis.

- Pre-use checks —  Some instruments require pre-use verification against a reference (e.g. tape measures against a master). Required in addition to formal calibration, not instead of.

# 5. Calibration Providers

Calibration is performed by:

- NATA-accredited providers —  Preferred for all M&TE. Calibration certificates show NATA endorsement.

- Other accredited providers —  Acceptable where NATA-accredited service unavailable. Equivalent mutual recognition accreditation required.

- Manufacturer service —  Acceptable for specialised equipment. Provider's accreditation must be documented.

- Internal calibration —  Only where internal facilities are accredited and traceability to national standards is documented.

Calibration certificates are reviewed on receipt — uncertainty figures, traceability and identification confirmed before the instrument is returned to service.

# 6. Use of M&TE

- M&TE must show clearly its calibration status — sticker showing due date or a portal lookup.

- M&TE that is out of calibration (overdue) must not be used for acceptance decisions.

- Users must check calibration status before using an instrument.

- Damaged or suspect M&TE must be taken out of service and the QA Inspector notified.

- Personal instruments used for company work are subject to this procedure on the same basis as company-owned equipment.

# 7. Out-of-Calibration Response

If an instrument is found to be out of calibration (drift detected at calibration, damage suspected, or calibration date passed):

- Remove the instrument from service immediately. Tag DO NOT USE.

- Notify QA Inspector and Quality Manager the same working day.

- Identify all measurements taken with the instrument since its last valid calibration.

- Assess whether nonconforming product may have been accepted based on suspect measurements.

- If product may have been released to the customer, raise an NCR (QMS-FRM-NCR-001) and CAR (QMS-FRM-CAR-001) — customer notification per QMS-PRO-002.

- Document the assessment regardless of outcome — the assessment is the audit evidence.

- Recalibrate or replace the instrument before returning to service.

- Where drift is significant, shorten the calibration interval for that instrument permanently.

# 8. Traceability

Calibration results must be traceable to national or international measurement standards through an unbroken chain of comparisons. Calibration certificates must reference:

- The standards used and their traceability.

- The accreditation body and accreditation number.

- The measurement uncertainty for each measurement.

- Identification of the instrument calibrated.

- Calibration date and conditions.

Where no national standard exists for a particular measurement, the basis of calibration is documented and approved by the Quality Manager.

# 9. Out-of-Service M&TE

M&TE that is damaged, suspect, awaiting recalibration, or no longer needed:

- Is tagged OUT OF SERVICE.

- Is removed from production areas to a quarantine location.

- Has its register status updated.

- Is repaired and recalibrated before return to service — never just recalibrated without addressing the cause.

- Is permanently disposed of where uneconomical to repair — register updated.

# 10. Records

| Calibration Register (QMS-REG-CAL-001) | Maintained by QA Inspector. Reviewed monthly. |
| --- | --- |
| Calibration certificates | Filed by QA Inspector. Retained life of equipment + 3 years. |
| Out-of-calibration assessments | Filed with NCR (if raised). Retained 7 years. |
| Internal calibration records | Where internal calibration is performed. Retained 7 years. |
| Pre-use check records | Where required. Retained 1 year. |

# 11. Related Documents

| Document No. | Title |
| --- | --- |
| QMS-PRO-003 | Control of Nonconforming Outputs |
| QMS-PRO-005 | Product and Service Release |
| QMS-REG-CAL-001 | Calibration Register |
| QMS-FRM-NCR-001 | Nonconformance Report Form |

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| QA Inspector |  |  |  |
| Quality Manager |  |  |  |
| Managing Director |  |  |  |

✎ HOW TO COMPLETE THIS DOCUMENT — DELETE THIS SECTION BEFORE PUBLISHING

Use these notes to tailor, complete and finalise this document before publication. Delete this section once done.

1. Complete the M **&****TE inventory in QMS-REG-CAL-001** Every instrument used for acceptance decisions must be in the register before this procedure can operate.

2. Confirm the NATA-accredited provider relationship Identify 1-2 NATA-accredited providers covering your typical instruments. Get a service agreement.

3. Set up due-date alerts Schedule alerts 60/30/14/7 days before calibration due. Portal can automate this.

4. Stock OUT OF SERVICE tags Have RED OUT OF SERVICE / DO NOT USE tags at every workshop bay and at the QA office.

5. Train all M **&****TE users** Every employee who uses an instrument for acceptance must know: check status before use, recognise damage, report immediately if out of cal.

6. Document personal-tool policy If employees use personal verniers etc, decide if they're acceptable for company work. If yes, they go in the register. If no, supply company tools.

7. Test the out-of-cal response Run a tabletop exercise — find an out-of-cal item and walk through the steps in Section 7. Identify any gaps in the response capability.

8. DELETE this guidance section before publishing Select the entire yellow-bordered block and delete.

QMS-PRO-006  •  Rev 2.0	CONTROLLED DOCUMENT — verify currency before use	Page  of